## Slide 1
Beam Instrumentation, Controls, Feedback and Operation Aspects
(Main Classification 6)
IPAC 2026 Student Programme 2026
Peter Forck
Gesellschaft für Schwerionenforschung (GSI), p.forck@gsi.de

## Slide 2
Demands on Beam Diagnostics and Control
Beam diagnostics is the ’sensory organs’ for a real beam in a real environment.
Beam control is the active influence on beam parameters by a person or machine.
Different demands lead to different installations:
Quick, non-destructive measurements leading to a single number or simple plots
Used as a check for online information.; reliable technology is required
Example: Beam position (=center-of-mass) by Beam Position Monitor BPM
Complex instruments for severe malfunctions, accelerator commissioning & development
The instrumentation might be destructive and complex
Example: Emittance determination, tune etc. measurement
General usage of beam instrumentation:
Monitoring of beam parameters for operation, beam alignment & accel. development
Instruments for automatic, active beam control mostly by a feedback system
Example: Closed orbit feedback at synchrotrons using position measurement by BPMs

## Slide 3
Demands on Beam Diagnostics
Non-invasive ( = ’non-intercepting’ or ‘non-destructive’) methods are preferred:
The beam is not influenced  at transfer line: the same beam at several locations
 at synchrotron: properties of the stored beam are monitored
The instrument is not destroyed due to high beam power
Instruments could be different for:
Transfer lines with single passage  synchrotrons with multi-passages
Electrons are mostly relativistic  protons are at the beginning non-relativistic
Remark:
Most instrumentation is installed outside of rf-cavities to prevent for signal disturbance
Beam diagnostics is the ’sensory organs’ for a real beam in a real environment.
Beam control is the active influence on beam parameters by a person or machine.

## Slide 4
Outline of the Tutorial
Part 1 of the lecture on Beam Position Monitors:
Beam position monitors for bunched beams
Position measurements and feedback system
Measurement of tune, chromaticity etc.
Part 2 of the lecture on transverse and longitudinal diagnostics:
Profile measurement methods
Transverse emittance measure, incl. tomography
Measurement of longitudinal parameters

## Slide 5
Outline for BPM measurements:
Signal generation and basic usage:
transfer impedance, button BPM
Electronics basics
Cavity BPM at FELs
Trajectory measurement, close orbit feedback
Tune and lattice function measurement
Measurements with BPMs

## Slide 6
Pick-Ups for bunched Beams
The image current at the beam pipe is monitored on a high frequency basis
i.e. the ac-part given by the bunched beam.
Beam Position Monitor BPM is the most frequently used instrument!
Relativistic transformation of an electric field:
A Beam Position Monitor is an non-destructive device for bunched beams.
It delivers information about the transverse center of the beam:
Trajectory: Position of an individual bunch within a transfer line or synchrotron
Closed orbit: Central orbit averaged over a period much longer than a betatron oscillation
Single bunch position: Determination of parameters like tune, chromaticity, -function
Remarks: - BPMs have a low cut-off frequency  dc-beam can’t be monitored
- The abbreviation BPM and pick-up PU are synonyms

## Slide 7
-
-
-
-
-
+
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
V
-
-
-
-
-
-
-
-
+
+
+
+
+
-
-
+
-
+
-
+
-
+
-
-
-
-
+
+
+
-
+
+
-
+
-
-
+
-
+
-
-
-
-
-
+
+
+
-
+
+
-
-
+
-
-
-
-
-
-
-
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
Animation by Rhodri Jones (CERN)
Principle of Signal Generation of a BPMs: off-center Beam
The image current at the wall is
monitored on a high frequency basis
i.e. ac-part given by the bunched beam.

## Slide 8
Model for Signal Treatment of capacitive BPMs
Equivalent circuit
The wall current is monitored by a plate or ring inserted in the beam pipe:
Ubeam
ground
CC
C
R
Ibeam
capacitive coupling
 image current
electronics elements
Uim
At a resistor R the voltage Uim from the image current is measured.
Goal: Connection from beam current to signal strength by transfer impedance
in frequency domain:
Result: i.e. complex function
geometry
stray capacitance
frequency response

## Slide 9
Principle of Position Determination by a BPM
The difference voltage between plates gives the beam’s center-of-mass
most frequent application
is called position sensitivity; sometimes the inverse is used
is a geometry dependent, non-linear function,
Units: =[%/mm], sometimes =[dB/mm] or =[mm].
It is at least:
from
Typical desired position resolution:
of beam width
Example: One turn = 4 bunches @ 35 MeV/u
Uright
Ureft

## Slide 10
2-dim Model for a Button BPM
‘Proximity effect’: larger signal for closer plate
Ideal 2-dim model: Cylindrical pipe  image current density
via ‘image charge method’ for ‘pencil’ beam:
Image current: Integration over finite button size:
,  = 0o


a
r
button
 =0

## Slide 11
Ideal 2-dim model: Non-linear behavior and hor-vert coupling:
Sensitivity converts signal to position
with = [%/mm] or [dB/mm]
i.e. is the derivative of the curve , here i.e. non-linear.
For this example: central part =7.4%/mm  mm
button
 real position
• measured position
Position Map
Horizontal plane
2-dim Model for a Button BPM


a
r
button


## Slide 12
Button BPM Realization
LINACs, e--synchrotrons: 100 MHz < frf < 3 GHz  bunch length  BPM length
 50  signal path to prevent reflections
Example: LHC-type inside cryostat:
24 mm, half aperture a = 25 mm, C = 8 pF
 fcut= 400 MHz, Zt = 1.3  above fcut
Ø24 mm

5cm
N-Connector
Button electrode
Ceramic vacuum insulation
Courtesy C. Boccard (CERN)

## Slide 13
Simulations for Button BPM at Synchrotron Light Sources
Result: - Non-linearity and xy-coupling occur in dependence of button size and position
- Can be corrected by polynomial interpolation for beams much smaller than chamber
real beam position x [mm]
y = 20 mm
y = 10 mm
y = 0 mm
20
Sx(center) = 7.8 %/mm
for |x|<5mm & y=0mm
Sy(center) = 7.2 %/mm
for |y|< 5mm & x=0mm
button 2
position x [mm]
30
button 1
button 3
button 4
30
position x [mm]
button 1
button 2
button 4
button 3
28 mm
72 mm
 7 mm
button 1
button 2
button 3
button 4
18 mm
y = 20 mm
y = 10 mm
y = 0 mm
1 mm
steps
1 mm steps
y
x
A.A. Nosych et al.,
IBIC’14
Optimization:
Horizontal distance
Size of buttons
Horizontal: 
Vertical: y  y
Example: Simulation for ALBA light source for 72 x 28 mm2 chamber

## Slide 14
Outline for BPM measurements:
Signal generation and basic usage:
transfer impedance, button BPM
Electronics basics:
schemes for signal processing
Cavity BPM at FELs
Trajectory measurement, close orbit feedback
Tune and lattice function measurement
Measurement with BPMs

## Slide 15
Broadband Signal Processing
Hybrid or transformer close to beam pipe for analog generation or
Attenuator/amplifier
Filter to get the wanted harmonics and to suppress stray signals
ADC: digitalization  followed by calculation of of
.
Advantage: Bunch-by-bunch observation possible, versatile post-processing possible
Disadvantage: Resolution down to  100 µm for large aperture, i.e.  0.1% of aperture,
resolution is worse than narrowband processing, see below
Challenge: Precise analog electronics with very low drift of amplification

## Slide 16
General Noise Consideration
1. Signal voltage given by:
2. Thermal noise voltage given by:
Example: GSI-LINAC with frf=36 MHz
Signal-to-noise is influenced by:
Input signal amplitude
Thermal noise from amplifiers etc.
Bandwidth
 Restriction of frequency width
as the power is
concentrated at harmonics

## Slide 17
Narrowband Processing for improved Signal-to-Noise
Narrowband processing, analog correspondence heterodyne receiver (AM-radio, spectrum analyzer)
Attenuator/amplifier
Mixing with accelerating frequency
 IF-output: signal with difference frequency
Bandpass filter of the mixed signal Rectifier: synchronous detector
ADC: digitalization → followed calculation of
Advantage: Spatial resolution about 100 time better than broadband processing
Disadvantage: No turn-by-turn diagnosis, due to mixing = ’long averaging time’
Digital
correspondence:
I/Q de-modulation
As today’s technology

## Slide 18
Outline for BPM measurements:
Signal generation and basic usage:
transfer impedance, button BPM
Electronics basics:
Schemes for signal processing
Cavity BPM at FELs:
information storage for short pulses
Trajectory measurement, close orbit feedback
Tune and lattice function measurement
Measurement with BPMs

## Slide 19
Flash, Hamburg
LINAC-based Light Sources:
Eelectron  1 … 18 GeV, coherent light from undulator, E < 1000 keV, temporally short pulse
4th Generation Light Sources & Beam Delivery
Single bunch duration < 100 fs:
Goal: Short bunches with high number of particles
short, intense laser pulses for electron generation
Requirement: Position stability  resolution < 1 µm
facc = 1.3 GHz
B. Steffen et al, DIPAC 2009, B. Steffen et al., Phys. Rev. AB 12, 032802 (2009)
2 = 110 fs

## Slide 20
TM01
TM11
TM02
U / V
frequency f [GHz]
monopole TM010
U  Ibeam
dipole TM011
U  Ibeamx
monopole TM020
Cavity BPM for FEL-LINACs: Principle
High resolution on t << 1 μs time scale can be achieved by excitation of a dipole mode
Depictive statement: ‘The position information is stored due to cavity mode excitation’
‘δ-excitation’ followed by
oscillation with Q  1000
and τ =2Q/2f  100 ns
 ‘0.1 μm resolution
within 100 ns‘
From D. Lipka,
DESY, Hamburg
For pill box the resonator modes given by geometry:
Monopole mode TM010 with f010
 maximum at beam center  strong excitation
Dipole mode TM011 with f011
 minimum at center  excitation by beam offset
Detection of dipole mode amplitude
voltage U
beam
x
pillbox cavity
antenna 1
antenna 2
monopole
TM010
dipole
TM110

## Slide 21
Suppression of mono-pole mode: waveguide only for dipole-mode transmission
Monopole Mode
Dipole Mode
Courtesy of D. Lipka,
DESY, Hamburg
Prototype BPM for ILC Final Focus:
Achieved resolution (i.e. 3 BPM relative) of 8.7 nm in a 6 × 12 mm beam pipe at ATF2 (KEK, Japan)
Remark: Separated, smaller cavity for monopole mode to have same frequency for normalization
Courtesy of D. Lipka,
DESY, Hamburg
wave guide
Monopole mode
Courtesy of D. Lipka & Y. Honda
Courtesy of D. Lipka and Y. Honda et al., Phys. Rev. Accel. Beams 11, 062801 (2008)
Cavity BPM: Suppression of monopole Mode
Example: fdipole,hori= 5.7 GHz, fdipole,vert= 6.4 GHz
Dipole mode
Wave-guide with cut-off frequency
:
blocks monopole-mode
transmits dipole-mode
Notes: 21

## Slide 22
Outline for BPM measurements:
Signal generation and basic usage:
transfer impedance, button BPM
Electronics basics:
Schemes for signal processing
Cavity BPM at FELs
information storage for short pulses
Trajectory measurement, close orbit feedback:
Beam stabilization methods via real-time control
Tune and lattice function measurement
Measurement with BPMs

## Slide 23
Trajectory:
The position delivered by an individual bunch within a transfer line or a synchrotron.
Main task: Control of matching (center and angle), first-turn diagnostics
Example: LHC injection 10/09/08 i.e. first day of operation !
Courtesy R. Jones (CERN)
-10
-10
horizontal
vertical
Monitor number (530 BPMs on 27 km)
10
-10
0
x [mm]
10
-10
0
y [mm]
Tune values at LHC: Qh = 64.3, Qv = 59.3
Trajectory Measurement with BPMs

## Slide 24
Close Orbit Measurement with BPMs
Orbit measurement:
Example: 12 beam positions at GSI-SIS during ramping from 8.6 to 500 MeV/u for Ar18+
feedback off
time [ms]
400
300
100
0
position [mm]
20
10
0
-10
200
position x(t) at 12 BPMs
Closed orbit:
Beam position averaged over many turns
(i.e. betatron oscillations).
The result is the basic tool for alignment & stabilization
Position resolution :
Required resol. depends on beam size :
Role of thumb:

## Slide 25
Close Orbit Feedback: Results
Orbit feedback:
Example: 12 beam positions at GSI-SIS during ramping from 8.6 to 500 MeV/u for Ar18+
Procedure:
1. Position digitalization of all 12 BPMs
2. Calculation of corrector setting on fast (FPGA-based) electronics
3. Submission to corrector magnets
4. New position measurement
 regulation time down to 10 ms
feedback off
time [ms]
400
300
100
0
position [mm]
20
10
0
-10
200
feedback on
time [ms]
0
100
300
400
position [mm]
20
10
0
-10
200
Rule of thumb:
Movement related to tune i.e. ‘natural oscillations by periodic focusing’
To determine the ‘sine-like’ oscillation 4 BPMs per oscillation are required
 4 BPMs per tune value (but detailed investigation required to determine the # of BPMs)
position x(t) at 12 BPMs

## Slide 26
Closed Orbit Feedback: Typical Noise Sources
Courtesy M. Böge (PSI) CAS 2008, N. Hubert (Soleil) DIAPC 2009
Beam movement:
Short term (min to 10 ms):
Traffic
Machine (crane) movements
Water & vacuum pumps
50 Hz main power net
Medium term (day to min):
Movement of chambers
due to heating by radiation
Day-night variation
Tide, moon cycle
Long term ( > days):
Ground settlement
Seasons, temperature variation
1
0.1
100
10
1000
Frequency [Hz]
Power spectral density [mm2/Hz]
0.01

## Slide 27
Close Orbit Feedback: BPMs and magnetic Corrector Hardware
Orbit feedback: Synchrotron light source  spatial stability of light beam
Example: SLS-Synchrotron at Villigen, Switzerland
Feedback loop:
1. Position from all BPMs
2. Calculation of corrector setting via
Orbit Response Matrix
3. Change of magnet setting
1.’ New positon measurement …..
 regulation time down to 10 ms
From M. Böge, PSI
Uncorrected orbit: typ.
Corrected orbit: typ. up to  100 Hz bandwidth!
Position from all BPMs
Calculation of corrector strength
Setting of correctors
feedback
Acc. optics
Rule of thumb: 4 BPMs per tune value (to determine the ‘sine-like’ betatron oscillation)

## Slide 28
Close Orbit Feedback: Mathematical Treatment
Linear beam optics lead to Orbit Response Matrix ORM:
Kick angle of single steerer transferres to an offset at BPM :
( transverse tune, phase advance, : beta functions )
For all steerers and all BPMs the
Orbit Response Matrix is defined as:
The matrix elements must be measured beforehand by
changing each steerer and observing the position at each BPM
Closed orbit feedback:
Measurement of the position at all BPMs,
Determine offset to respect to set value
Calculation of new steerer kick for correction
determination of pseudo-inverse ORM
frequently used method: Singular Value Decomposition
ORM matrix
position
at BPM
angle
at steerer
pseudo-inverse
ORM
position
at BPM
angle
at steerer

## Slide 29
Close Orbit Feedback: Controller Principle
A feedback is modelled by Control Theory!
Visualization for a control loop:
+
-
Controller
System
Sensor
Reference orbit
Deviation
Corrector
settings
Corrected
closed orbit
Newly measured
closed orbit
Orbit
feedback
Synchrotron
BPM
System output
Measurement output
Mathematical description for multi-input multi-output (MIMO) system:
Synchrotron optics: Orbit Response Matrix
System reaction: Model for power supplies, beam dynamics, tune spread etc.
BPM: Accuracy and noise
Using Laplace Transformation and frequency domain visualization for dynamic behavior
Reference

Error
Correction
Courtesy Guenther Rehm (BESSY), IBIC 2019 & CAS 2026

## Slide 30
Close Orbit Feedback: System Bandwidth
Reaction depends on bandwidth and latency:
Performance measure by cross over frequency and latency :
Beam damping below by anti-phase correction of disturbance
Beam excitation above by in-phase excitation with respect to disturbance
Large bandwidth required to enable fast damping
Courtesy Guenther Rehm (BESSY), IBIC 2019 & CAS 2026
Lower crossover
 lower bandwidth
 slower damping
damping region
excitation region
excitation region
damping region
Higher latency
 lower bandwidth
 slower damping
Slow feedback: time step of orbit correction much lower than system delay
Fast feedback: time step of orbit correction comparable to system delay
Frequency [Hz]
10
100
1000
104
Magnitude

## Slide 31
Close Orbit Feedback: Actual Challenges
Courtesy Guenther Rehm (BESSY), IBIC 2019 & CAS 2026
Lower crossover
 lower bandwidth
 slower damping
damping region
excitation region
excitation region
damping region
Higher latency
 lower bandwidth
 slower damping
Slow feedback: time step of orbit correction much lower than system delay
Fast feedback: time step of orbit correction comparable to system delay
Frequency [Hz]
10
100
1000
104
Magnitude
Actual challenges (mainly for 4th generation light sources):
Measurement accuracy for BPMs
High bandwidth regulation
High bandwidth power supplies
Model inclusion of alignment errors
 Technical improvements
Magnets with non-linear field contributions
Space charge and impedance contributions
Linear approx. by ORM might be insufficient
Proposal of Machine Learning
for non-linear problems

## Slide 32
Outline for BPM measurements:
Signal generation and basic usage:
transfer impedance, button BPM
Electronics basics:
Schemes for signal processing
Cavity BPM at FELs
information storage for short pulses
Trajectory measurement, close orbit feedback:
beam stabilization methods via real-time control
Tune and lattice function measurement
Measurement with BPMs

## Slide 33
Graphics by R. Singh, GSI
Tune Measurement: General Considerations
Beam particle’s in-coherent motion  center-of-mass stays constant
Coherent excitations are required for the detection by a BPM
Excitation of all particles by rf  coherent motion
 center-of-mass variation turn-by-turn i.e. center acts as one macro-particle

## Slide 34
Tune Measurement: The Kick-Method in Time Domain
The beam is excited to
coherent betatron oscillation:
 Beam position measured
each revolution (’turn-by-turn’)
 Fourier Trans. gives non-integer tune q.
Short kick compared to revolution.
Here: N = 200 turn  q > 0.003
(tune spreads can be q  0.001!)
The de-coherence time limits the resolution:
N non-zero samples
 General limit of discrete FFT:
Decay is caused by de-phasing,
not by decreasing
single particle amplitude.

## Slide 35
Tune Measurement: Frequency Chirp Measurement
SIS-synchrotron: A wide scan with
both sidebands at h=25th-harmonics:
From the position of the sidebands q = 0.306 is determined.
Advantage: High resolution, possibility of tune feedback by phase detection
Disadvantage: Long sweep time (up to several seconds).
Alternative: Excitation by band-limited noise with beam ‘picks’ the correct frequency
A detailed scan for the lower sideband
 beam acts like a driven oscillator:
Principle: Slowly scan of the excitation frequency  beam acts as driven oscillator!
(sometimes refer to as Beam Transfer Function BTF or frequency chirp measurement)


## Slide 36
Chromaticity ξ: Change of tune for off-momentum particle
Two step measurement procedure:
1. Change of momentum p by detuned rf-frequency
2. Excitation of coherent betatron oscillations
and tune measurement
(kick-method, BTF, noise excitation):
Chromaticity Measurement from Closed Orbit Data
Plot of ΔQ/Q as a function of Δp/p
 slope is dispersion ξ.
Example: Measurement at LEP:
From M Minty, F. Zimmermann,
Measurement and Control of charged Particle Beam,
Springer Verlag 2003

## Slide 37
Actual Challenges for BPMs
Actual challenges for BPMs and COFB:
 Position resolution requirement at 4th generation circular light sources : about 1 µm (!)
Low drift analog and digital electronics
Solution keywords: Pilot tone or channel multiplexing
Fast signal processing providing high bandwidth & low latency
Solution keywords: Parallel signal processing on powerful FPGAs
Improved accelerator model
Solution keywords: Machine Learning to include non-linearities
Actual challenges for tune etc. measurement:
Low excitation strength
Solution keywords: Phase-sensitive tracking (PLL-tracking)
Possible tune & chromaticity feedback
Solution keywords: Under development
Related talks at IPAC’26 within MC6:
Laura Torino (ALBA), Overview of the beam position monitors development for the next generation light sources

## Slide 38
Outline for profile measurements:
Scintillation screens:
emission of light, universal usage, limited dynamic range
Optical Transition Radiation:
light emission due to crossing material boundary, mainly for relativistic beams
Ionization Profile Monitor and Beam Induced Fluorescence
Synchrotron Light Monitors
Measurement of Beam Profile

## Slide 39
Scintillation Screen
Scintillation: Particle’s energy loss in matter causes emission of light
the most direct way of profile observation as used from the early days on!
Pneumatic drive with Ø70 mm screen:
Flange
200 mm
& window
Screen
 70 mm
beam
Pneumatic
drive
Camera
 1 m

## Slide 40
Alumina: Al2O3
CsI:Tl
Chromox: Al2O3:Cr
P43
YAG:Ce
Quartz
Quartz:Ce
ZrO2:Mg
Very different light yield i.e. photons per ion‘s energy loss
Different wavelength of emitted light
Light output from various Scintillating Screens
Example: Color CCD camera: Images at different particle intensities determined for U at 300 MeV/u
Type | Name | Material | Activ. | Max. λ | Decay
Cera- mics | Chromox | Al2O3 | Cr | 700nm |  10ms
 | Alumina | Al2O3 | Non | 380nm |  10ns
Crystal | YAG:Ce | Y3Al5O12 | Ce | 550nm | 200ns
 | LYSO | Lu1.8Y.2SiO5 | Ce | 420nm | 40ns
Powder of gains Ø10μm on glass | P43 | Gd2O3S | Tb | 545nm | 1ms
 | P46 | Y3Al5O12 | Ce | 530nm | 300ns
 | P47 | Y2SiO5 | Ce&Tb | 400nm | 100ns

## Slide 41
Optical Transition Radiation: Depictive Description
Optical Transition Radiation OTR for a single charge e:
Assuming a charge e approaches an ideal conducting boundary e.g. metal foil:
Image charge is created by electric field with dipole type field pattern
Field distribution depends on velocity  and Lorentz factor  due to relativistic trans. field increase
Penetration of charge through surface within t < 10 fs: sudden change of source distribution
Emission of radiation with dipole characteristic
perfect
metal
vacuum
charge e
velocity 
E-field
pattern
dipole type
image
charge -e
velocity -
perfect
metal
vacuum
charge e
velocity 
image
charge -e
velocity -
perfect
metal
vacuum
charge e
inside
metal
dipole
radiation

 max 1/
Physics: sudden change charge distribution
rearrangement of sources  radiation
Other physical interpretation: Impedance mismatch at boundary leads to radiation


## Slide 42
Technical Realization of Optical Transition Radiation OTR
OTR is emitted by charged particle passage through a material boundary.
Insertion of thin Al-foil under 45o
Observation of low light by CCD.
Photon distribution:
within a solid angle d and
Wavelength interval begin to end
Detection: Optical 400 nm <  < 800 nm
Larger signal for relativistic beam  1
Low divergence for  1  large signal
Well suited for e- beams
p-beam used for Ekin 10 GeV   10

## Slide 43
Outline:
Scintillation screens:
emission of light, universal usage, limited dynamic range
Optical Transition Radiation:
light emission due to crossing material boundary, mainly for relativistic beams
Ionization Profile Monitor and Beam Induced Fluorescence:
secondary particle detection from interaction beam-residual gas
Synchrotron Light Monitors
Measurement of Beam Profile

## Slide 44
Interaction between Residual Gas and the Beams
Physics: Energy loss of ions in gas dE/dx
Profile determination from residual gas
Ionization:
in average roughly  100 eV/ionization
Excitation followed by photon emission:
in average roughly  300 eV/event e.g. N2
However, depends strongly on gas molecule
e-
e-
res. gas atom
e-
e-
e-
e-
e-
beam
beam pipe
e-
ionization
e-
excitation
1H
12C
40Ar
238U
Energy loss in mbar N2 by SRIM
source
LINAC, cycl.
synchr.

## Slide 45
Ionization Profile Monitor at GSI Synchrotron
Non-destructive device for proton synchrotron:
Beam ionizes the residual gas by electronic stopping
Gas ions or e- accelerated by E -field 1 kV/cm
Spatial resolved single particle detection
Realization at GSI synchrotron:
One monitor per plane
R
voltage divider
+ 6 kV
+ 5 kV
+ 4 kV
+ 2 kV
+ 3 kV
+ 1 kV
0 kV
MCP
anode
position readout
HV electrode
ion
e.g. H+
beam
Abbreviation: Ionization Profile Monitor IPM or
at CERN Beam Gas Ionization BGI Monitor

## Slide 46
Ionization Profile Monitor Realization
The realization for the heavy ion storage ring ESR at GSI:
Horizontal camera
Horizontal IPM:
E-field box
MCP
IPM support
& UV lamp
Ø250 mm
beam
E-field separation disks
View port Ø150 mm
electrodes
Insertion 650 mm
175mm
Vertical IPM
Vertical camera
Realization at GSI synchrotron:
One monitor per plane

## Slide 47
Large beam power  Non-intercepting method:
Beam Induced Fluorescence BIF
N2 + Ion  (N2+)*+ Ion + e-  N2+ + γ + Ion + e-
With single photon detection scheme
390 nm <  < 470 nm blue wavelength
non-destructive, compact installation.
Example: Installation of hor&vert. BIF Monitor:
Beam Induced Fluorescence for intense Profiles
Fluorescence gas:
Either from residual gas or extended gas bump
Or intersecting gas jet

## Slide 48
Outline:
Scintillation screens:
emission of light, universal usage, limited dynamic range
Optical Transition Radiation:
light emission due to crossing material boundary, mainly for relativistic beams
Ionization Profile Monitor and Beam Induced Fluorescence:
secondary particle detection from interaction beam-residual gas
Synchrotron Light Monitors
standard for circular light sources, visible or X-ray range
Measurement of Beam Profile

## Slide 49
Synchrotron Light Monitor
An electron bent (i.e. accelerated) by a dipole magnet emit synchrotron light.
This light is emitted
into a cone of
opening 2/ in lab-frame.
Well suited for rel. e-
For protons:
Only for energies Ekin > 100 GeV
The light is focused to a camera
Advantage:
Signal anyhow available!
orbit of electrons
orbit of electrons
dp/dt
dp/dt
radiation field
radiation field
power: P  4/  2
Rest frame of electron:
Laboratory frame:

## Slide 50
Synchrotron Light Monitor overcoming Diffraction Limit
Problem: The diffraction limit is µm for a typical case
Direct optical observation:   100 µm
 Relative simple installation, but resolution insufficient for modern light sources
Direct X-ray observation with shorter  :   10 µm
 Medium complex but expensive installation,
installed at most synchr. light sources referred to as pinhole camera
Interference optical observation:   1 µm
 Complex installation, seldom installed; basis: blurring of interference pattern

## Slide 51
Actual Challenges for Profile Measurement
Actual challenges for profile measurement:
High resolution and dynamic range
Solution keywords: Material choice, complicated optical setup,
sensitive electronics and/or optics
Non-invasive methods
Reason: high beam power at hadron acc., multi-location observation
Solution keywords: Machine Learning based image reconstruction
Related talks at IPAC’26 within MC6:
Renjun Yang (CSNS), Achieve a record dynamic range of halo diagnostics with a novel fluorescence wire concept
Carsten Welsch (University of Liverpool), Progress in Data-Driven Beam Diagnostics within LIV.INNO
 Conclusion

## Slide 52
Outline for transverse emittance diagnostics:
Quadrupole strength variation and position measurement
emittance from several profile measurement and beam optical calculation
Measurement of transverse Emittance

## Slide 53
Definition of Coordinates and basic Equations
The basic vector is 6 dimensional:
The transformation of a single particle from a location s0 to s1 is given by the Transfer Matrix R:
The transformation of a the envelope from a location s0 to s1 is given by the
Beam Matrix  :
6-dim Beam Matrix with decoupled hor., vert. and long. plane:
Horizontal
beam matrix:
Beam width for
the three
coordinates:
horizontal
vertical
longitudinal
hor.-long. coupling
 10 values

## Slide 54
Definition of Beam Matrix: 6-dim versus 2-dim Matrices
Generally, the beam matrix
contains all 6-dim variances
(i.e. correlations):
For linear forces, the equation is valid for all dimensions
For linear forces, the 6-dim emittance is preserved:
Remarks:
A 2-dim matrix was used in textbooks providing simpler formulas
Non-linear forces are e.g.:  Sextupole, octupole etc. i.e. higher-order magnetic fields
 Non-linear contribution by longitudinal rf as it is a sine-wave
 Space charge forces for intense beams or cooled beams
 Impedance contributions
is symmetric matrix
due to etc.
has 21 independent quantities if all directions are coupled.

## Slide 55
Emittance Measurement by Quadrupole Variation
From a profile determination, the emittance can be calculated via linear transformation,
if a well known and constant distribution (e.g. Gaussian) is assumed.
The basis is beam matrix transformation
Measurement of beam width
With the drift matrix the transfer is
Transformation of the beam matrix
Task: Calculation of matrix
at entrance i.e. all three elements

## Slide 56
Measurement of transverse Emittance under ideal Conditions
Example: Square of the beam width at
ELETTRA 100 MeV e- Linac, YAG:Ce screen:
Using the ‘thin lens approximation’ i.e. the quadrupole has a focal length of f:
Measurement of matrix-element from matrices (s1,Ki) = R(Ki )  (s0)  RT (Ki )
G. Penco (ELETTRA) et al., EPAC’08
Focusing strength K= 1/f [m-1]
For completeness: The relevant formulas
A fit delivers the beam matrix elements ij(s0)
Assumptions:
Well aligned beam, no steering
‘Regular’ phase space distribution
No emittance blow-up due to space charge
Improved methods:
Based on e.g. tomographic reconstruction
Focus
at screen
Focus
upstream
Focus
downstream
 Emittance conclusion

## Slide 57
Tomographic Reconstruction
original object
Iterative projection
& back-projection
after sufficient
iterations
1st back-
projection
Tomography: Generation of a n-dim image
from many lower-dim projections
Medical application: 2-dim reconstruction using sufficient 1-dim projections
Filtered back projection:
Iterative process by redistributing
the 2-dim image and considering the
differences to the previous iteration step.
Courtesy S. Hancock et al.

## Slide 58
original object
Iterative projection
& back-projection
after sufficient
iterations
1st back-
projection
Most often used:
FBP: Filtered Back-Projection
ART: Algebraic Reconstruction Technique
Other methods:
MENT: Maximum Entropy
Image recognition by Machine Learning
Application at accelerators:
Transverse & longitudinal
emittance reconstruction in transfer lines
Longitudinal bunch evolution in synchr.
Tomography: Generation of a n-dim image
from many lower-dim projections
Medical application: 2-dim reconstruction using sufficient 1-dim projections
Tomographic Reconstruction
Courtesy S. Hancock et al.

## Slide 59
Experiments at ALICE dedicated transfer line Daresbury Laboratory (UK):
Electron LINAC ALICE with MeV
in the frame of EMMA FFAG (up to year 2016), then CLARA
Dedicated line with typ. 600 betatron phase advance
between Yag:Ce screens (0.75 m distance each)
Several quadrupoles for flexible beam setting
for versatile experiments
Exemplary Results of 2-dim Emittance Reconstruction
K.M. Hock, A. Wolski et al., NIM A 753, 38 (2014), M.G. Ibison et al., JINST 7, P04016 (2012), IPAC 2013
Typical 2-d trans. profile
@ YAG:Ce#1
1.5 m
YAG:Ce screens 1 to 3
5 cm
hor. x
vert. y
 8 mm
Typically 150 quad-setting Quad 7
(200… 1700 betatron phase)
& profile measurements screen 1
tomography by FBP
Phase space at entrance of Quad 7
for low current (20 pC/bunch)

## Slide 60
Experiments at ALICE dedicated transfer line Daresbury Laboratory (UK):
Beam current effect detected
Exemplary Results of 2-dim Emittance Reconstruction
Increase of current
Phase space in front of Quad 7
for low current (20 pC/bunch)
Phase space in front of Quad 7
for high current (80 pC/bunch)
Related to
Space charge defocusing upstream Quad 7
Scan of Quad 7 &
measure at screen 1
Consistency check:
Comparison of measured beam profile and reconstructed profile at reconstruction location

## Slide 61
Actual Challenges for Emittance Measurement
Actual challenges for emittance measurement:
High resolution and dynamic range
Solution keywords: Material choice, sensitive electronics and/or optics
Reconstruction algorithms
Solution keywords: Tomography with optimized ‘traditional’ methods,
Machine Learning based image reconstruction
Related talks at IPAC’26 within MC6:
Sarah Geffroy (Université Paris-Saclay, CNRS/IN2P3, IJCLab),
Emittance measurement of the CERN antimatter beam for the GBAR experiment
Mayu Wada (The University of Tokyo),
Development of a new ultraslow muon beam diagnostic system for the J-PARC muon g-2/EDM experiment
Related talks at IPAC’26, MC6 on Machine Learning:
Eito Iwai (RIKEN), ML-driven automated tuning of SACLA XFEL: progress and future
Ysabella Kassandra Ong (INFN), AI and machine learning techniques for LNL accelerators
Thorsten Hellert (LBNL),
Bridging the Gap Between Control Systems and Natural Language: A Framework for Semantic Channel Finding
Fuhao Ji (SLAC), Intelligent scientific facility R&D: AI/ML developments at SLAC MeV-UED
 Conclusion

## Slide 62
Measurement of longitudinal parameter:
Bunch length measurement at
FEL-LINACs
Circular synchrotron light facilities
Longitudinal  transverse correspondences:
position relative to rf  transverse center-of-mass
bunch structure in time  transverse profile
momentum or energy spread  transverse divergence
longitudinal emittance  transverse emittance.
Measurement of longitudinal Parameters

## Slide 63
Longitudinal Emittance by linear Transformation using a Buncher
Longitudinal focusing:
Variation of the bunch shape by a rf-buncher
 components 5 and 6 from 6-dim phase-space
Transversal correspondence: Quadrupole variation
 Transfer matrix of buncher & drift:
with focal length
 Variation of buncher amplitude
 different bunch width at monitor location
LINAC
RFQ
source
Bunch Shape Monitor
B.
Buncher
BSM

## Slide 64
Bunch Length Measure by electro-optical spectral Decoding
LINAC
LINAC
EO monitor
Undulator
Bunch compressor
Courtesy S.P.Jamison et al., EPAC 2006
For Free Electron Lasers  bunch length below 1 ps is achieved
Below the resolution of streak camera
Short laser pulses with t  10 fs and electro-optical modulator
Electro optical modulator: Birefringent, rotation angle depends on external electric field
Relativistic bunch: transverse elecitc field carries the time information
Spectral decoding
Optical stretcher: Separation of colors by different path length
 different colors at different time  single-shot observation

## Slide 65
Hardware of a spectral-decoded EOSD Scanning Setup
B. Steffen et al, DIPAC 2009
B. Steffen et al., Phys. Rev. AB 12, 032802 (2009)
crystal
beam
Example: Bunch length at FLASH
100 fs bunch duration = 30 µm length!
2 = 110 fs
vacuum flange
S. Funkner et al., arXiv1912.01323 (2019)
Laser: Commercial Ti:Sa or Yb-fibre,
10 fs duration, near IR,
Crystal: GaP or ZnTe, 100 µm thickness
 Long. conclusion

## Slide 66
Results of tomographic Reconstruction at a Synchrotron II
Bunches from 500 turns at CERN PS and phase space for time slice of shortest bunch,
measured with a Wall Current Monitor WCM:
Fast rise of rf-amplitude for ‘bunch-rotation’:
Bunch is compressed in time
Filamentation due to amplitude- dependent synchrotron frequency
extraction
injection
synchrotron
Current trans.
FCT or WCM

## Slide 67
Bunch Length Measurement for relativistic Electrons
Electron bunches are too short (t < 100 ps) to be covered by the bandwidth of
pick-ups (f < 3 GHz  trise > 100 ps) for structure determination.
 Time resolved observation of synchr. light with a streak camera: Resolution  1 ps.
injection
extraction
Streak camera
e- synchrotron
acc.
freq. frf
rf cavity
Scheme of a streak camera
Input optics
Streak tube
CMOS
camera
 60 cm

## Slide 68
Bunch observation by Streak Camera
 Longitudinal mismatched injection leads to coherent synch. oscillations, coupled bunch motion
Example: Bunch shape oscillations at Elettra, Trieste, MHz :
From M. Ferianis, CAS’07
The streak camera delivers a fast scan in horizontal direction (here 441 ps full scale)
and a slower scan in vertical direction (here 9 ns).
injection
extraction
Streak camera
e- synchrotron
acc.
freq. frf
rf cavity

## Slide 69
Actual Challenges for Bunch Shape Measurements
Actual challenges for bunch shape measurement:
High resolution and dynamic range
Solution keywords: FEL-LINACs costly methods such as Transverse Deflecting Cavity
Reconstruction algorithms
Solution keywords: Tomography with optimized ‘traditional’ methods,
Machine Learning based image reconstruction
For FELs: Synchronization of cavities and experiential trigger
Solution keywords: Electro-optical signal transmission, ultra-accurate electronics
Related talks at IPAC’26 within MC6:
Frank Ludwig (DESY), From femtosecond to attosecond RF field control

## Slide 70
Conclusion for Instrumentation and Control
Beam instrumentation is required to control the ‘real-world’ beam correctly
Beam Position Monitors:
Non-invasive, versatile application, sensor for feedback
Profile detectors:
Various detector types, non-invasive preferred
Emittance determination:
Reconstruction by beam optics and tomography, non-linear solutions e.g.by ML
Bunch length determination:
Various methods, partly complex or costly instruments, tomography
Thank you for your attention
Do you have questions or comments?