## Slide 1: First Experience with the UE17 APPLE3 Afterburner at FLASH2
Markus Tischer on behalf of DESY’s insertion device group
ID25 Workshop, Taipei, 31.05.2025

## Slide 2: Topics
APPLE3 Afterburner for FLASH2 
is a prototype for 6 FLASH2020+ Seeding IDs
Undulator Group FS-US
  Philip Eckoldt
  Mona Gehlot
  Kathrin Götze
  Paul Neumann
  Andreas Schöps
Mechanical Design Department (ZM1)
  Hilmar Bienert
  Hakan Bolat
Production Planning (ZM2)
Photon Science Machine Shop (FS-BT)
FLASH Team (MFL*, FS-FLASH)
* thanks to Juliane Rönsch-Schulenburg for various slides
CoWorkers
Major Contributions by
  Markus Tischer
  Pavel Vagin
  Thorsten Vielitz
  Sumera Yamin
  Torsten Ramm
  Peter Talkovski
Introduction
Commissioning studies and user experiments
Operation experience
Changes UE17  UE35

## Slide 3
315 m
5 MeV
150 MeV
Bunch Compressors
550 MeV
Accelerating Structures
RF Stations
Lasers
RF Gun
Fixed Gap Undulators
Xseed & X-Ray
Photon Diagnostics
THz
FLASH1
FLASH2
Albert Einstein
Kai Siegbahn
Variable Gap Undulators
FLASH3
FEL Experiments
TESLA type superconducting accelerating modules 1.3 GHz
FLASH1 fixed gap undulators
3rd harmonic sc module 3.9 GHz
Albert Einstein Hall
FLASH2 variable gap undulators
Kai Siegbahn Hall
350MeV -
1350 MeV
Normal conducting 1.3 GHz RF gun
Cs2Te cathode / 2 injector lasers
Extraction to FLASH2
Laser heater
FLASH
Notes: FLASH is a 300 m long single-pass high-gain soft x-ray FEL. It is based on a superconducting accelerator using the TESLA/XFEL technology. An here you see a layout of FLASH how it looked like about a year ago. The electron bunches generated in a normal conducting RF gun and are accelerated by the superconducting accelerator modules.

## Slide 4
315 m
5 MeV
150 MeV
Bunch Compressors
550 MeV
Accelerating Structures
RF Stations
Lasers
RF Gun
Fixed Gap Undulators
Xseed & X-Ray
Photon Diagnostics
THz
FLASH1
FLASH2
Albert Einstein
Kai Siegbahn
Variable Gap Undulators
FLASH3
FEL Experiments
FLASH2 variable gap undulators
Kai Siegbahn Hall
=> Up to 5000 bunches per second
TESLA type superconducting accelerating modules 1.3 GHz
FLASH1 fixed gap undulators
3rd harmonic sc module 3.9 GHz
Albert Einstein Hall
Laser heater
350MeV -
1350 MeV
FLASH
Notes: This allows the generation of several thousand bunches per second. The bunches are generated in burst with a repletion rate of 10 Hz.

## Slide 5
315 m
5 MeV
150 MeV
Bunch Compressors
550 MeV
Lasers
RF Gun
Fixed Gap Undulators
Xseed & X-Ray
Photon Diagnostics
THz
FLASH1
FLASH2
Albert Einstein
Kai Siegbahn
Variable Gap Undulators
FLASH3
FEL Experiments
FLASH2 variable gap undulators
Kai Siegbahn Hall
FLASH1 fixed gap undulators
Albert Einstein Hall
Laser heater
Δp
Δp
t
t
350MeV -
1350 MeV
=> Up to 5000 bunches per second
FLASH
Notes: A laser heater system minimizes the microbunching instability growth by controlling the uncorrelated energy spread of the beam

## Slide 6
315 m
5 MeV
150 MeV
Bunch Compressors
550 MeV
Lasers
RF Gun
Fixed Gap Undulators
Xseed & X-Ray
Photon Diagnostics
THz
FLASH1
FLASH2
Albert Einstein
Kai Siegbahn
Variable Gap Undulators
FLASH3
FLASH1 fixed gap undulators
Albert Einstein Hall
Laser heater
Δp
Δp
t
t
FLASH2
350MeV -
1350 MeV
SASE FLASH2:
Photon energy range: 390 - 14 eV
Pulse duration: 1.3-200 fs
Min. spectral width: 0.5%
Afterburner:
Polarization: variable
Photon energy range: 890 - 80 eV
=> Up to 5000 bunches per second
FLASH

## Slide 7
ext. seeded
315 m
5 MeV
150 MeV
Bunch Compressors
550 MeV
Lasers
RF Gun
FLASH1
FLASH2
Albert Einstein
Kai Siegbahn
Variable Gap Undulators
FLASH3
Laser heater
Δp
Δp
TDS
THz
t
t
350MeV -
1350 MeV
Simulation
Seeded FLASH1:
Photon energy range: 310 – 20.6 eV
Pulse duration: 11- 40 fs
Min. spectral width: 0.05%
Afterburner:
Polarization: variable
Photon energy range: 890 - 80 eV
SASE FLASH2:
Photon energy range: 390 - 14 eV
Pulse duration: 1.3-200 fs
Min. spectral width: 0.5%
=> Up to 5000 bunches per second
Future FLASH

## Slide 8
5 MeV
150 MeV
Bunch Compressors
550 MeV
Lasers
RF Gun
FLASH1
FLASH2
Albert Einstein
Kai Siegbahn
Variable Gap Undulators
FLASH3
350MeV -
1350 MeV
Afterburner:
Polarization: variable
Photon energy range: 890 - 80 eV
SASE FLASH2:
Photon energy range: 390 - 14 eV
Pulse duration: 1.3-200 fs
Min. spectral width: 0.5%
FLASH2
FLASH2 Afterburner – UE17 APPLE3 with Force Compensation | 
Period length | 17.5 mm
K value | 0.92
B0 | 0.56 T
Gap | 7 mm
N | 137
Length | 2.5 m
Polarization | any
Wavelength range | 1.33 . 1.77 nm
Energy range | 890 – 700 eV
@1.35GeV

## Slide 9: UE17 Afterburner Operation Overview
Sep’23 UE17 installation at FLASH2 (circ. 6mm vac.chamber already in operation since Aug’22)
Oct-Nov’23 2 Commissioning shifts: Beam transport, gap and shift dependent kick errors
          No issues found
Jan-Mar’24 3 shifts to establish 3rd harm radiation
          Slow progress in preparing an intense low wavelength beam with high 3rd harm. bunching
          3rd harm radiation established with S/B ratio ~100
          Optimization of the driving SASE beam (rev. taper, laser heater)
Spring’24 1st user beamtime
Early Jun’24 Begin of FLASH shutdown
          Ending Aug 3rd ‘25

## Slide 10
Main SASE undulator : U32 planar
Helical afterburner : UE17 APPLE3
Need to get rid of powerful linearly polarized radiation from the main undulator
helical afterburner
2.5 m
planar undulator (saturation)
30 m
reverse-tapered planar undulator (K increases)
Fully micro-bunched electron beam but strongly suppressed radiation power at the exit of reverse-tapered planar undulator
The beam radiates at full power in the helical afterburner tuned to 3rd harmonic
Reverse Taper Experiments
Suppression of radiation of main undulators

## Slide 11
Reverse Undulator Tapering for Polarization Control and Background-Free Harmonic Production in XFELs: Results from FLASH,- Schneidmiller & Yurkov; DOI: 10.18429/JACoW-FEL2017-MOP032
Bunching factor at saturation can be the same as in the case of a non-tapered undulator,
Saturation length increases moderately while the saturation power is suppressed by orders of magnitude
17 nm
Previous Reverse Taper Experiments
Test with FLASH2 main (SASE) undulators
Notes: In 2016 reverse taper was applied to the 10 undulator segments; the gap of the 11th and 12th segments was scanned. Power ratio of 200 was obtained. For a helical afterburner it would be larger by a factor of 2.
The figure shows the photon pulse energy measured after the respective undulator in FLASH2. Undulators 1 to 10 were set to reverse tapering while the final two act as an afterburner, leading to an increase in pulse energy by a factor of 200.

## Slide 12
Afterburner tuned to 3rd harmonic
Afterburner detuned
Mono
Grating
Screen
Undulators
Helical
Afterburner
New double monochromator Beamline FL23 used setup to distinguish fundamental and 3rd harmonic
Successful Suppression of Linear Polarization
Suppression of main undulator radiation while obtaining polarized afterburner radiation

## Slide 13
80
70
60
50
40
30
Single Pulse energy (a.u.)
time
LH off
LH optimized
The output power of the afterburner could be doubled by optimizing the longitudinal phase space distribution using the laser heater
Focusing into the afterburner helps to increase the output power especially for short wavelength with long gain length.
Influence of Electron Beam Quality
Laser heater (LH)

## Slide 14
scanned
wavelength
~ 3 µJ at 3 nm
amplification by a factor of 80
~ 2 µJ
amplification by > 100
Different Working Points
80 eV – 860 eV demonstrated
Notes: amplification by a factor of 80

## Slide 15: Recent Experimental Results at FLASH Beamline FL23
Element-specific probe of ultra-fast magnetic properties
Spectral characterization of the helical afterburner radiation, benchmarking of experimental setup
M. Pavelka et al., Struct. Dyn. 12, 024303 (2025)

## Slide 16: Polarization Diagnostics by Photoelectron TOF Spectrometers
Angular distribution of photoelectron emission of rare gas excitations
P1= 0.98
P1= 0.99
ABU shift = 0.0
ABU shift = +/- 0.5
ABU shift = +/- 0.25
P1= 0.08
P1= 0.98
P1= 0.06
Ee = 880 MeV
hv = 104.6 eV / 313.8 eV
Test of different analysis methods for a robust fast analysis tool, ~2% accuracy
Kr 3d
Courtesy of M. Braune, proceed. SRI2024

## Slide 17: Polarization Diagnostics by Photoelectron TOF Spectrometers
Angular distribution of photoelectron emission of rare gas excitations
P1= 0.98
P1= 0.99
ABU shift = 0.0
ABU shift = +/- 0.5
ABU shift = +/- 0.25
P1= 0.08
P1= 0.98
P1= 0.06
Ee = 880 MeV
hv = 104.6 eV / 313.8 eV
Test of different analysis methods for a robust fast analysis tool, ~2% accuracy
More detailed error analysis (E3rd = 318.8 eV)
Negligible unpolarized component ~  (10-3)
Kr 3d
S. Marotzke et al., Struct. Dyn. 12, 034301 (2025);

## Slide 18: Polarization Diagnostics by Photoelectron TOF Spectrometers
Gap scans
3rd harmonic signal intensity change along afterburner gap variation  Verify the optimal gap value
Value of linear Stokes parameter P1 along afterburner gap variation
Courtesy of M. Braune, proceed. SRI2024

## Slide 19: XAFS / XMCD Measurements at 3d L-Edges
Benchmarking by comparison of reference sample measurements
S.Marotzke et al., Struct. Dyn. 12, 034301 (2025);

## Slide 20: Radiation Dose Measurements
Diagnostic undulator at FLASH2
  Based on TLDs
  Total charge (2016-2024) ~7800mC
  Demagnetization rate ~1% / 40 kGy
Average dose at Sase IDs ~ 4-10 Sv/100mC
Online dose measurement system put to operation
  Based on (shielded) RadFET sensors
  Temporal structure relates to TLDs
  Dose values need to be scaled
Field loss vs dose
Field loss vs time
US
DS
Center

## Slide 21: Design “Review” UE17
“Assorted” selection of items
Gap mechanics - 4-axes version, useful for alignment of ID and vacuum chamber
Shift drive: - some elastic deformation, observed by a sin-like mismatch of lin. vs motor-encoders
Shift girders - poor initial alignment, needed mech. shimming throughout - requires more precise definition of CMM protocols
Keeper / subgirder interface ok
Keeper design - single period keeper ok, but could be simplified
Force compensation works very well; larger mech. moments expected & ok
Magnets - glued AB-pairs without problems
Single keeper measurements: - indispensable for sorting - mech. initialization to be improved
Magnetic sorting, measurements and tuning - demanding, ok, room for improvement

## Slide 22: Design Changes from UE17 to UE35
Gap mechanics - 4-axes version unchanged
Shift drive: - 4-axes as before - incoupling of shift motion moved closer to the gap to reduce torque and deformations of drive train - movable shift girders thickened
Shift rails - unchanged, but qualified another supplier
Period length - 17.5  35 mm
Keeper design - mostly unchanged - gap of comp. magnets enlarged by~0.2mm to ease alignment of vacuum chamber
Magnet blocks - unchanged, glued AB-pairs

## Slide 23: Construction Status of 6 APPLE3 UE35 for FLASH2020+
Impressions…
Single keeper measurements ~80% completed
3 IDs completely constructed
Tuning nearly completed for 2 IDs 3rd device started
Construction of remaining IDs is ongoing

## Slide 24: Thank you

## Slide 25: Backup

## Slide 26: Vacuum Chamber: Resistive Wall & Surface Roughness
Resistive wall wake :
  20 keV E-loss & E-spread for a bunch of 200 pC, 100 fs and 2.5 m chamber length
  ok for Al with (inner) R < 600 nm, and oxide layer < 10 nm
Courtesy of J. Zemella, M. Vogt
Extruded aluminum chamber
  7x7mm2 cross-section, Dm 6mm inner bore
  Length ~2.5m
  Support by 0.5mm thick blades in a groove
  Adjustment by movers

## Slide 27: Magnetic Performance
(simulated)
Different Gap-dependence of vertical (By) and horizontal (Bz) field components
Condition for circularly polarized radiation deviates from the nominal values for both gap and shift
To be implemented in ID-controls by a 2-dimensional table based on magnetic and x-ray measurements at a later stage

## Slide 28
Gap- and Shift-Dependence of K-value