## Slide 1: Beam Tests of a Permanent Magnet Medical Accelerator Arc from 10-250MeV

## Slide 2: Hadron Therapy Accelerator
Dejan Trbojevic proposes to accelerate protons in 10-250MeV range with a rapid cycle (500Hz+) for FLASH hadron therapy
  Many cycles within ~100ms allow many energy (depth) slices to be irradiated as well as being quick enough to trigger the biological FLASH effect
Need a fixed-field (FFA) ring with fixed tunes
  But needs to fit in the hospital, so smaller than a scaling FFA  nonlinear non-scaling FFA

## Slide 3: Nonlinear Non-Scaling FFAs
Linear field non-scaling FFAs have cell tunes that vary with momentum
This tune dependence can be flattened by adding non-linear components to the magnet fields (although not necessarily for an unlimited momentum range)
Quadrupole affects tune level, sextupole affects dQ/dp, octupole affects d2Q/dp2, etc.

## Slide 4: ‘FLASH’ Proton Therapy facility: Stony Brook University Hospital – Radiation Oncology
6. D. Trbojevic, Title: “Non-scaling fixed field alternating gradient permanent magnet cancer therapy accelerator”, patent number: US 9661737 B2, Date of the patent: May 23, 2017 (belongs to the DOE). https://patentimages.storage.googleapis.com/42/5e/92/f7da1cf617d6e3/US9661737.pdf

## Slide 5: Facility Parameters
Parameter | Value | Unit
Injection Ek | 10-30 | MeV
Extraction Ek range | 65-250 | MeV
Arc radius | 3.5 | m
Machine cycle rate | 500-800 | Hz
Energy gain per turn | 30-45 | kV
Treatment charge | 60 | nC
Treatment time | 75-250 | ms

## Slide 6
Matching of the Multiple Energy Orbits to the Straight Section
3.5m arc radius in current design

## Slide 7: FFA Arc Cell Tunes
After iterative correction in Muon1, starting from an initial BMAD model, tunes are stable from 10MeV – 250MeV and nearly constant.
(Wiggly near low energy end because rigidity changes fast there).
Cell tune ranges:
Qx = 0.1608 ± 0.0004
Qy = 0.0738 ± 0.0011

## Slide 8: Nonlinear Magnet Field Profiles
Maximum fields equalised at 1.86T rather than 1.99T
“original” = BMAD model

## Slide 9: Test Beamline Simulated Orbits
10MeV
250MeV
HF
HF
F
F
F
D
D
D
D
10cm grid
Angle 5.625° per cell, 4 cells total = 22.5° (1/8 of an 180° arc)
Half-F magnets on each end  parallel orbits
Cell length 34.4cm, 4 cells total = 1.374m

## Slide 10: Permanent Magnet Cross-Sections
F
D
1cm grid

## Slide 11: Permanent Magnet Parameters
Parameter | F magnet | D magnet | Unit
Material | NdFeB |  | 
Grade | N42EH |  | 
Remnant field Br | 1.30 |  | T
Area | 69.90 | 93.88 | cm2
Full aperture | 64.7 × 12 | 48.1 × 12 | mm
Good field region x | ±26.3 | ±18.0 | mm
Maximum field | 1.85 | 1.86 | T

## Slide 12: Magnets Ordered from SABR Enterprises, LLC.

## Slide 13: Orbits within Magnet Models

## Slide 14: Test Beamline with Mounting

## Slide 15: Magnet Construction by Vendor
SABR Enterprises, LLC. in Massachusetts

## Slide 16: Beamline Assembly

## Slide 17: Magnet Apertures

## Slide 18: Fieldmap of one magnet

## Slide 19: Measured Integrated Field
Corrected via fit for Hall probe displacements and tilt angles, <1% error

## Slide 20: Camera and Beam Screen Plates
These were attached to both ends of the beamline.
4 cameras and 4 screens in total.

## Slide 21: Experiment at BNL beamlines
NSRL
  Protons 50-250MeV
  1 day 2 hours
Tandem
  Deuterons 12.9-28MeV
  Equivalent to protons 10-50MeV
    Via rigidity but also window/air energy loss calculation
    Tandem proton max energy is 28.75MeV
  1 day

## Slide 22: Installation (Tandem)
Different beam energies marked on a ruler for alignment.

## Slide 23: Calibration and 1s Ellipse Fit

## Slide 24: Dimmest Beam Fit
Also background frames were subtracted

## Slide 25: NSRL 250MeV protons
Beam in
Beam out

## Slide 26: NSRL 200MeV protons
Beam in
Beam out

## Slide 27: NSRL 150MeV protons
Beam in
Beam out
Obscured by stray light on screen 3

## Slide 28: NSRL 100MeV protons
Beam in
Beam out

## Slide 29: NSRL 70MeV protons
Beam in
Beam out

## Slide 30: NSRL 50MeV protons
Beam in
Beam out

## Slide 31: NSRL All Energies
Beam in
Beam out

## Slide 32: Tandem 28MeV d ~ 50MeV p
Beam in
Beam out

## Slide 33: Tandem 24.2MeV d ~ 40MeV p
Beam in
Beam out

## Slide 34: Tandem 20MeV d ~ 30MeV p
Beam in
Beam out

## Slide 35: Tandem 16MeV d ~ 20MeV p
Beam in
Beam out

## Slide 36: Tandem 14.3MeV d ~ 15MeV p
Beam in
Beam out

## Slide 37: Tandem 12.9MeV d ~ 10MeV p
Beam in
Beam out
Beam may have stopped in air before screen 4

## Slide 38: Tandem All Energies
Beam in
Beam out

## Slide 39: All 12 Ellipses
Beam in
Beam out

## Slide 40: Screen Alignment Errors
This affected the Tandem run, so the screens were resurveyed and the offsets subtracted
Screen 1
Screen 2
Screen 3
Screen 4

## Slide 41: Corrected Centroid X Positions
Beam in
Beam out
Location tracks with energy as expected
Alignment was “by hand”: 3-8mm input errors

## Slide 42: Filling the Dynamic Aperture?
At lower energies, the output beam gets very large and starts to assume a crescent shape
Simulation
Experiment

## Slide 43: Experiment Conclusion
Beams at all energies were transmitted through the 4-cell test girder (22.5° or 1/8 arc)
  Protons at NSRL and deuterons at Tandem
10-250MeV is a factor of 5.3× in momentum
  Very large for a non-scaling FFA, especially one with a flattened tune
  250MeV is highest ever energy in any NS-FFA line
Beam ought to propagate several cells around a ring at the start of commissioning
  (even without magnet shimming/correction here)

## Slide 44: BACKUP

## Slide 45: Fast Cycling Synchrotron Cycles
t(ms)
Kinetic
Energy
250
(MeV)
16.5-30
One synchrotron cycle
Ex.1
Ex.3
Ex.2
Injection
Energy
Second synchrotron cycle
0.8ms < t < 1.85ms
1.3kHz< f < 540 Hz
RF manipulation
Extraction #2
Time of bunch flight 405-550 ns
At the injection porch time of 926 ms there are 1680 turns. 15 turns for 15 Cyclotron bunches and 1650 turns for RF manipulations
Injection
Porch
If the RF kicks are 30 kV, during the slip-stacking merge a difference in the time of flight for two 2 energies is 773ps. If the batches are 15 m apart it would take 370 turns to get them aligned. Next step would be to merge into one bucket by snaping the new RF voltage on.
Extraction kicker wave forms

## Slide 46: Fast Cycling Acceleration with Ferrite cavities
 | Choices for injector cyclotron |  |  |  | Units
 | ACSI | GE | TR-19 | IBA | 
Energy of Injector Cyclotron | 14 | 16.5 | 19 | 30 | MeV
Cyclotron current | 400 | 200 | 400 | 1200 | mA
Treatment charge | 60 | 60 | 60 | 60 | nC
Injection turns | 4 | 4 | 4 | 2 | #
RF acceleration per turn | 29.70 | 29.70 | 29.70 | 44.55 | kV
Calculation | Treatments Options |  |  |  | 
Turn duration | 585.74 | 540.6 | 504.8 | 405.1 | ns
Charge per turn | 0.234 | 0.108 | 0.151 | 0.486 | nC
Charge per injection | 0.937 | 0.432 | 0.606 | 0.972 | nC
Machine cycles | 65 | 139 | 100 | 62 | #
Machine cycle rate | 540 | 540 | 540 | 809.95 | Hz
Outputs |  |  |  |  | 
Treatment time | 120 | 257 | 185 | 76.54 | ms
By Stephen Brooks

## Slide 47: Peak Fields vs. Energy
Orange line is field at magnet longitudinal centre, which is different than at exit (orbit in fringe area) for D magnet.
D

## Slide 48: Orbits and Fields in Muon1

## Slide 49: Lattice Converted from BMAD
Multipoles dipole…dodeca
Three cells for field wrap
Angle 5.625° per cell, 64 cells per turn. Cell length 34.4cm, R=3.5m.
This worked well in BMAD, fixed stable tunes from 10MeV – 250MeV
These slides look a lot like last year’s, but Dejan’s multipole numbers have changed!

## Slide 50: BMAD and Muon1 Magnet Models
BMAD has a zero-length end field model
Muon1 has a Maxwellian fringe field model with finite length (bit more realistic)
  But you have to choose the length yourself!
  Only way to do this sort of field in BMAD is to generate fieldmaps, so there’s a difference
We also changed from sectors to rectangles but that turned out to be a small effect

## Slide 51: Field Profiles (BMAD)

## Slide 52: BMAD Lattice in Muon1
Vertical tune is unstable at low energies
  On previous lattice version, this was true no matter what Muon1 fringe field length I chose
MeV
Cell tune

## Slide 53: This cell can be “easily” fixed using the Algorithm I used before
Runge-Kutta 4th order tracking step
Loop to get trajectory in cell
Finite difference to get transfer matrix
  Also gives tunes if orbit is closed
Iterate (Newton) to find closed orbit
Loop over all FFA energies
Finite difference parameters to get response matrix of tune functions to multipole changes
Iterate (optimiser) to find fixed-tune lattice
Muon1
Muon1 cell optics mode
Muon1 FFA optics mode
Fixed-tune FFA optimiser
(parallel)
Fourier coefficient

## Slide 54: Wait, Fourier coefficients?
Representing the transverse midplane field dependence By(x,0) as a Fourier series made optimisation work better than multipoles
  Suspect this is because sin,cos(nx) are orthogonal whereas polynomials are highly correlated
Used the form
By(x,0) = c0 + Sn=1..8 snsin(nkx) + cncos(nkx)
…where 2p/k is about twice the orbit excursion

## Slide 56: Magnet Sim. Field Error Graph (F)
3.6×10-4 relative error

## Slide 57: Magnet Sim. Field Error Graph (D)
1.6×10-4 relative error

## Slide 58: Permanent Magnet Designs (F)
cm grid
NdFeB material
N42EH grade
Br=1.3T
69.90cm2 area
64.7 x 12mm aperture
Good field X±26.3mm
1.85T max field

## Slide 59: Permanent Magnet Designs (D)
NdFeB material
N42EH grade
Br=1.3T
93.88cm2 area
48.1 x 12mm aperture
Good field X±18.0mm
1.86T max field
cm grid

## Slide 60: Magnet Holders Machined

## Slide 61: Magnet Holders Machined

## Slide 62: End Plates Machined

## Slide 63: Raw Integrated Field Data

## Slide 64: Measured Gradients

## Slide 65: Attempt at Shimming
Shimming is possible in theory but more engineering is needed to keep wires securely in place in aperture, so magnets were not shimmed for this experiment.

## Slide 66: Test Beamline, 4 cells
Half-F magnets on each end
Angle 5.625° per cell,
4 cells = 22.5°
(1/8 of an 180° arc)
Cell length 34.4cm,
4 cells = 1.374m

## Slide 67: Centroid X Positions
Beam in
Beam out

## Slide 68: Corrected Centroid X Positions
Beam in
Beam out

## Slide 69: Centroid Y Positions
Known Y shift from reassembly of camera screens between NSRL and Tandem
Beam in
Beam out

## Slide 70: Attempt at Transfer Matrix Element
Not much time, but some energies were tried offset horizontally from the nominal entrance position by a few mm
  Often the entrance position of the previous energy was used with the next energy
dxout/dxin ~= (change in average X centroid position on output screens)/(change in average X centroid on input screens)

## Slide 71: Transfer Matrix Mxx Element
4Qx,cell = 0.644  cos(fx) = -0.618
Agrees at high energy, changes at low energy
  Could be injecting off the closed orbit  sextupole feed down
(Kind of surprised we got any valid data for this at all)

## Slide 72: Tracking Offset Trajectories
±10mm

## Slide 73: Matrix Element from Tracking
Looks sensitive to injection position especially at low energy
±10mm

## Slide 74: Transfer Matrix Mxx Element Again
With the corrected centroids, use the difference of the beam position to the central orbit (according to survey) to derive a matrix element between input and output
  No longer calculated around an offset orbit!
  Looks more reliably negative, but noisier
  1mm centroid errors make large error bars on this

## Slide 75: Magnetic Efficiency
Circumference or field factor of 2.29× relative to ideal synchrotron

## Slide 76: Two Stage Machine
1.63× relative to ideal synchrotron
Momentum ratio 5.31=3.54*1.40, split at 135MeV k.e.

## Slide 77: Two Stage Machine Conclusion
The two stage machine reduces fields by a factor of 1.4, or same field at lower radius
  One stage FFA: 3.5m radius (7m diameter)
  Two stage FFA: 2.5m radius (5m diameter)
  Ideal synchrotron: 1.53m radius (3.06m diameter)
    But difficult to ramp the magnets at 100s of Hz
Quick design: just cut out pieces of the original machine’s momentum range and scale them!